TESTS AND ADJUSTMENT PROCEDURES

FOR THE DEMODULATION BOARD.

These tests are performed with the anti-alias filters and the compression filters.

Before tests, place the delay lines in a centered position.

- Equipment necessary for procedures:

  - Spectrum/Network analyzer HP 35665 A or equivalent.

  - Preamplifier.

  - 2 Stanford research systems DS345 generators.

  - Oscilloscope.

  - General power supply DC +/-15V & +/-6V.

  - Halogen lamp with DC 12V, 3A power supply.

- Functionalities:

Verify that the general power supply is not in short circuit.

Verify the numeric function “REARM” is working well.

- BLU demodulation tests:

Using one DS345, perform the LO signal: Sinus; 0dBm; 6,25MHz.

Using the other DS345, perform the RF signal: Sinus; 1Vpp; 6,251MHz.

Synchronize the two generators: connect the “10MHz” output from the LO generator to the RF generator “timebase” input.

Using the oscilloscope, verify the 0 and 90 outputs: verify the output signals are 1kHz Sinus.

Note the amplitudes and calculate the two gains for the two channels.

- Phase adjustment:

Using the network analyzer, perform the function: 0 output/ 90 output.

Measure the phase difference and adjust it to 90, using the two delay line on the board.

- Noise measurements:

5-1 Demodulation board noise.

Connect a 50 resistor at the RF input.

Using the spectral analyzer, measure the output noise for the two channels (0, 90) for f=100Hz and 1kHz.

5-1 Global noise.

Connect the preamplifier to the RF input.

Measure the DC current using a self at the preamplifier output.

Perform the desired IDC value adjusting the halogen lamp current.

 Note the noise values for IDC=0 and IDC=20mA.

Subtract the Av amplifier noise if it is not negligible and calculate the ratio.

A last noise measurement can be performed using the LO board for the LO signal.