TESTS AND ADJUSTMENT PROCEDURES

FOR THE PHOTODIODE PREAMPLIFIER.

- Equipment necessary for procedures:

  - Spectrum/Network analyzer AGILENT 4395A or equivalent.

  - Low noise measurement amplifier with DC +/-5V power supply: Av=37,5dB , en =1,5nV/√Hz.

  - Halogen lamp with DC 12V, 3A power supply.

  - Demodulation board for preamplifier power supply.

- Frequency adjustment and transfer impedance measurement:

  - Connection:

First, remove the screw from the self L3.

Plug the definitive photodiode, protected from light.

The analyzer must be used in a network analyzer mode.

Set the display in a A/R configuration to perform the transfer function (gain & phase).

Frequency sweep from 5 MHz to 15MHz.

Source amplitude 0dBm.

Input range (A&R): 0dB.

IF BW = 1kHz.

- First, calibrate the Network analyzer with the power splitter and the cables:

source

R input

A input

-6dB

analyzer

50

50

50

- Then perform the transfer function:

-6dB

50

Source

R input

A input

analyzer

50

50

50

10k

preamplifier

Test

Out

    - Connect the test strap inside the preamplifier.

    - Adjust L1 to bring the notch frequency to 12,54MHz.

    - Adjust L2 to bring the peak frequency to 6,27MHz.

Adjust accurately the peak with the marker max. function.

    - Measure the peak gain (GdB) and notch gain.

Transfert impedance calculation:

Req≈10000×Glin

with Glin the linear gain Glin=10GdB20

 Note the Req value and verify that the level difference between the peak and the notch is at least 60dB.

 DO NOT USE A METALIC SCREW DRIVER FOR THE ADJUSTMENTS !!

 DO NOT FORGET TO REMOVE THE TEST STRAP !!

- Noise measurement.

The analyzer must be used in a spectral analyzer mode.

Center the frequency sweep at 6,27MHz and adjust the span (about 1MHz).

Input range: -20dB.

Average ON, average number = 50.

Measure the DC current using a self at the preamplifier output.

Perform the desired IDC value adjusting the halogen lamp current.

Source

A input

B input

analyzer

50

50

50

preamplifier

Test

Out

A

self

Av= 37,5dB

 Note the noise values for IDC=0 and IDC=20mA.

Subtract the Av amplifier noise if it is not negligible and calculate the ratio.